Contents

- Before use …………………………………………1

- Nomenclature ………………………………………2

- Assemblage ………………………………………….5

- Operation ……………………………………………7

- Configuration ………………………………………13

- Technical  parameter………………………………15

- Troubleshooting ……………………………………17

1 Before use

1-1 NOTICE

- Microscope ought to be placed in a dry and clean place. Do not expose the microscope in the sun directly. Avoid high temperature and violent vibration.

- As microscope is a precision instrument, handle with care, avoiding impact or abrupt movement during transportation.

- To keep the image clear, do not leave fingerprints or stains on the surfaces of the lens.

- Never turn the left and right focusing knob in the adverse direction at the same time, otherwise the microscope will be damaged.

- Hold the camera with one hand for fearing of falling when you take the films out of the big camera.

1-2 MAINTENANCE

1) All lenses must be kept clean. Fine dust on the surface of the lens should be blown off with hand blower or wiped off gently with a soft lens tissue; Fingerprints or oil marked on it should be wiped off with a tissue moistened with a small amount of xylene or a 3:7 mixture of alcohol and ether.

2) Never use the organic solution to clean the other surface (especially the plastic surfaces). If necessary, please choose the neutral detergent.

3) Do not take the microscope apart for fearing that it is damaged.

4) After using, cover the microscope with the dust-cover provided and store it in a dry and clean place free from moisture to prevent rust.

- To keep the performance of the microscope, please check it periodically. The detail can be gotten from the agent nearby.

2 Nomenclature

2-1 SZM-45B2

Eyepiece shield

Eyepiece

Diopter adjusting ring

Head

Illumination device

Zoom control knob

Focusing

Lamp lock-screw

Glass stage

Clip

Base

Top variable light control

Fuse

Bottom variable light control

Switch

2-2 SZMT2

Diopter adjusting ring  Eyepiece shield

Zoom control knob

Eyepiece

Head lock-screw

Head

Focusing knob

Glass stage

Base                                           Clip

2-2 SZM-45V2ZM

Eyepiece shield

Diopter adjusting ring

Eyepiece

Focusing knob

Illumination device

Digtal head

Lamp lock-screw

Base

Glass stage

Clip

3Assemblage

3-1 SZM-45B2

Eyepiece shield

Eyepiece

Head

Head lock-screw                                                   Illumination device

Clips

Glass stage

Base

3-2 SZM-45T2+SZM-PH/SZM-CTV

PK-mount

Photo eyepiece

Eyepiece shield

Eyepiece

Head

C-mount

Illumination device

Clip

Glass stage

Base

4 Operation

4-1 Use the glass stage

1) Press the glass stage on the sunken place then the other side of the glass stage will be lifted. (Fig.1)

Fig.1

4-2 Adjust the degree of tightness of the focusing arm.

1) If you want to adjust degree of tightness of the focusing arm, you     can hold one of the focusing knobs and turn another one to attain a suitable position. The degree of tightness relies on the direction to be turned. The clockwise direction is tight, otherwise, is loose.

2) The suitable position of tightness can make the adjustment more comfortable and prevent the focusing bracket from slipping down by its weight during the observation. (Fig.2)

Fig.2

4-3 Set the specimen slide

1) Set the specimen on the center of stage plate. If necessary, clamp the slide with the clips.

2) Turn on the light.

4-4 Adjust the specimen slide

1) Turn the zoom control knob to the maximum magnification.

2) Turn the diopter adjusting rings to the zero.

3) Observe the specimen through the right eyepiece and make the                                                                                                        jkj image clear by turning the focusing knob.

Fig.3             4) Rotate the zoom control knob to the minimum magnification.

5) Observe the specimen through the right eyepiece and make the image clear by turning the right diopter adjusting ring② .(Fig.3)

6) Redo the step(1),(3),(4)and (5) till the right adjusting ring is more precise.

7) Do the step (4) and make the image clear which is observed through the left eyepiece by turning the left diopter adjusting ring ①. (Fig.3)

4-5 Adjust the interpupillary distance

1) Adjust the prism housing along the direction of arrowhead of the

Fig.4 till the observation is comfortable.

4-6 Use Eyepiece shields

Fig.4

1) For user who does not wear glasses, hold the diopter-adjusting ring to prevent them from rotating and turn the eyepiece till the eyepiece shields fit the observer well.

2) For user who wears glasses, take the eyepiece shields off before

observation

4-7 Mount and Remove the Optional Eyepiece Micrometer

1) Turn and remove the mounting ring② from the eyepiece.(Fig.5)

2) Clean the eyepiece micrometer① and mount it to the mounting

ring with the inscription side downward.

3) Gently twist the mounting ring with the eyepiece micrometer into

the eyepiece till tightening② securely.

4) To remove the eyepiece micrometer, take down the mounting ring③ by twisting and take out of the micrometer, and then wrap it.

Fig.5

4-8 Install the illumination device

1)Insert the illumination device① in the bracket with the protrudent

side toward the lock-screw② and tighten the lock-screw. (Fig.6)

2) Put the plug into the socket of the pillar stand③.

Fig.6

4-9 Choose the optical system

1) You can alternate the binocular observation and video capture by

pushing or pulling the pole. You can attain binocular observation by

pushing the pole inside, or attain video capture by pulling it outside.

No matter what optical system is chosen, push or pull the pole

thoroughly.

Fig.7

4-10 Mount the photo eyepiece and the PK-mount              adapter

- Put the photo eyepieces socket of the tri-ocular.

- Connect the PK-mount adapter with the photo eyepiece, and then tighten the lock-screw. (Fig.8)

Fig.8

| Fig.9 Fig.10 Fig.11 Fig.12 | 4-11 Adjust the CTV - Adjust the CTV to a suitable position by rotating C-mount. Note: The range of the adjustment: 1~2mm in general.(Fig.9) 4-12 Connect the Digital head With the Monitor or TV set 1) Plug one end of the PVA cable into the socket of the digital head. (Fig.10) 2) Plug the C-VIDEO or S-VIDEO of the PVA cable into the correct socket of Monitor (TV set). 3) Connect the 12V DC power with the power socket of the PVA cable. 4-13 Appear the image on the Monitor or TV 1) Connect the power supply and then turn on the Monitor or TV. 2) For the monitor, the connect sign model must be chosen (C-video or S-video) and for TV, the channel must be set to the video channel. 3) Pull the pole out and adjust the focusing knob and then the image will appear on the screen clearly. 4-14 Connect with the computer 1) Plug one end of the PVA cable into the socket of the digital head. 2) Plug one of the C-VIDEO or S-VIDEO into the A/D board. 3) Plug the USB of the A/D board into the USB socket of the computer. (Fig.11) 4) If your computer is mounted capture card, you can connect the C-VIDEO or S-VIDEO with the computer directly. 5) Connect the 12V DC power with the power socket of the PVA cable. 4-15 Appear the image on the computer 1) Turn on the power supply and let the computer work. 2) Install the software and the driver of the A/D board. (If they have been installed, this step can be omitted. 3) Double click the icon of the software, and then the video window will appear. You can set the size of the window according to your linking 4) Draw out the pole and adjust the focusing knob, and then the image will appear on the computer screen clearly. 5) If no image or the image without color, it may be because the model of the input signal does not match the output signal of CCD or the model of C-VIDEO/S-VIDEO is not correct. The detail of operation refers to《Software operation manual》. 4-16 Appear the image on the computer and the Monitor synchronously 1) Do step 4-12 and step 4-14 to connect the computer and the Monitor. 2) Operate step 4-13 and step 4-15, we can make the image appear on the computer and Monitor at the same time. 4-17 Adjust the image - Put the base, stand and digital head correctly, then fix the lock-screw tightly. - Put the object on the base stage. - Observe the object through the eyepiece and adjust the focusing knob to make the image of the object clearly. - Move the digital head or the object gently to adjust the image agreeing with observer. 4-18 Brief instruction for the software - The program design of the software is up to date, and the Chinese/English interface can berth powerful delineation bar which be used much conveniently and rapidly. You can finish most of analyze work only to click the mouse. 2) Can afford many powerful area choosing tools which can analyse any area your linking at will, such as adjusting hue and image, dealing with mathematical morphology, image matching, texture analyse, character identify and so on. 3) Geometry character measuring function, automatically analyzing function such as slightness body, grain body, line body and so on. The outcome can be kept in data and can be made into chart and so forth. 4-19 Use the white balance 1) The CCD has auto white balance when the white balance switch is on“ON”. - Please put the switch on “ON” in general. Let the switch be “OFF” only in special, for example, observing the red cell, otherwise the color of red cell will be adjusted into white. - If you want to observe another single color, please let the switch be“ON”again when you finish the observation, and put the switch on “OFF” again after auto balance, or the color of the image will be distortion.(Fig.12) |
| --- | --- |

4-20 Adjust the brightness of the bottom light

1) Turn the adjustable light knob① according to the sign marked on

the base, along the clockwise the brightness will be added,

otherwise it will be weakened. (Fig.13)

Fig.13

4-21 Replace the lamps

1) Press the stage on the sunken place then the other side will be

lifted. (Fig.14)

2) Take the lamp out of the jack.

3) Put a new lamp into the jack thoroughly.

Fig.14             4) Recover the stage plate. (Fig.15)

Note: ① Before replacing the lamps, turn off the power first.

② Avoid violence while the lamp is plugged into the jack.

Fig.15

4-22 Replace the fuse

1) Screw the fuse tube out with a screwdriver and then pull the fuse

out of the tube①.

2) Renew the fuse and mount it in an adverse way. (Fig.16)

Fig.16

5 Configuration

5-1 SZM series configuration

| Configuration |  | Model |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Parts | Specification | SZM45B1 | SZM45B2 | SZM45B3 | SZM45T1 | SZM45T2 | SZM45T3 | SZM45V2 | SZM45P2 |
| Eyepieces | SZMEWh10X20 | O | O | O | O | O | O | O | O |
|  | SZMEWh15X15 |  |  |  |  |  |  |  |  |
|  | SZMEHWh20X10 |  |  |  |  |  |  |  |  |
| Binocular | SZM7045 | O | O | O |  |  |  |  |  |
| Tri-ocular | SZM7045TR |  |  |  | O | O | O |  |  |
| Digital head | SZM7045V |  |  |  |  |  |  | O | O |
| A/D board | W-A/D |  |  |  |  |  |  |  | O |
| PAV cable | W-RVB-PAV |  |  |  |  |  |  | O | O |
| Power adaptor | W-BW230-12-4.8 |  |  |  |  |  |  | O | O |
| Software | Pholib3.0 |  |  |  |  |  |  |  | O |
| Auxiliary objective | SZMAO0.5/165mm |  |  |  |  |  |  |  |  |
|  | SZMAO1.5/45mm |  |  |  |  |  |  |  |  |
|  | SZMAO2/30mm |  |  |  |  |  |  |  |  |
| Focusing arm | SZMA1 | O | O | O | O | O | O | O | O |
| Stand | SZMST1 | O |  |  | O |  |  |  |  |
|  | SZMST2 |  | O |  |  | O |  | O | O |
|  | SZMST3 |  |  | O |  |  | O |  |  |
|  | SZSTL1 |  |  |  |  |  |  |  |  |
|  | SZSTL2 |  |  |  |  |  |  |  |  |
| Transformer | SZT1 |  |  |  |  |  |  |  |  |
| Epi-illuminator | SZML1 |  | O | O |  | O | O | O | O |
| Hold for illuminator | SZFH1 |  |  |  |  |  |  |  |  |
|  | SZPD1 |  |  |  |  |  |  |  |  |
| Photo device | SZMPH |  |  |  |  |  |  |  |  |
| TV adapter | SZMCTV |  |  |  |  |  |  |  |  |
| Gem clamp | S/ST-GC |  |  |  |  |  |  |  |  |
| Dark field condenser Light | ST-30-2L-HJ-01 |  |  |  |  |  |  |  |  |
| Ring fluorescence | SZRL | O |  |  | O |  |  |  |  |
| Box | Inside foam Outside carton | O | O | O | O | O | O | O | O |

Note: The items marked“O”included and others for option

SZM series configuration

6Technical parameter

6-1 SZM7045/SZM7045TR

| Eyepiece | Standard configuration |  | Auxiliary objectives |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 0.5X |  | 1.5X |  | 2X |  |
|  | Working distance100mm |  | Working distance165mm |  | Working distance 45mm |  | Working distance 30mm |  |
|  | Magnification | Field of view | Magnification | Field of view | Magnification | Field of view | Magnification | Field of view |
| 10X/20 | 7X | 28.6 | 3.5X | 57.1 | 10.5X | 19 | 14X | 14.3 |
|  | 45X | 4.4 | 22.5X | 8.9 | 67.5X | 3 | 90X | 2.2 |
| 15X/15 | 10.5X | 21.4 | 5.25X | 42.8 | 15.75X | 14.3 | 21X | 10.7 |
|  | 67.5X | 3.3 | 33.75X | 6.7 | 101.25X | 2.2 | 135X | 1.7 |
| 20X/10 | 14X | 14.3 | 7X | 28.6 | 21X | 9.5 | 28X | 7.1 |
|  | 90X | 2.2 | 45X | 4.4 | 135X | 1.5 | 180X | 1.1 |

6-2 The base electronic specification of SZM series

| Model Parts |  | SZMST1 | SZMST2 | SZMST3 |
| --- | --- | --- | --- | --- |
| Power supply |  | No | 220V-50Hz、110V-50/60Hz | 220V-50Hz、110V-50/60Hz |
| Transformer |  | No | Input: 220/110VAC Output: 12V DC/45W | Input: 220/110VAC Output: 12V DC/45W |
| Illuminator | Top light | No | 12V/15W halogen lamp | 12V/15W halogen lamp |
|  | Bottom light |  | 12V/15W halogen lamp | 220/110V、7W fluorescent lamp |

6-3 CCD specification

| CCD Model Specification | NTSC | PAL |
| --- | --- | --- |
| Pick-up Device | 1/3 ″ | 1/3 ″ |
| Vertical Resolution | 450 TV line | 450 TV line |
| Number of Pixels | （H）768*（V）494 | （H）712*（V）582 |
| Scanning System | 525 lines, 60Field/Second | 625lines，50 Field/Second |
| White Balance | Can be switched between auto white balance and hand white balance |  |
| Back light compensation | Auto | Auto |
| Signal/Noise ratio | More than 46db | More than 46db |
| Gamma Characteristic | 0.45 | 0.45 |
| Minimum illumination | 3 Lux | 3 Lux |
| Input voltage | 12V DC（9V-14V） | 12V DC（9V-14V） |
| Power consumption | 1.85 or less | 1.85W or less |

- Working distance is fixed regardless of the magnification factor.

- Total mag.= Zoom mag. X Eyepiece mag. X Auxiliary objective mag.

Field number of eyepiece

Diameter of field of view（mm）=

Zoom mag.X Auxiliary objective mag.

- Photo adaptor mag.= Zoom mag.( ×Auxiliary objective mag.)×Eyepiece mag.

- TV adaptor mag.=Zoom mag.（Xauxiliary objective mag.）X C-mount TV adaptor middle

- Field of video view is 83%

- Total video magnification range is 18∽117

7 Troubleshooting

The performance of the microscope can’t be made fully because of unfamiliar using, this table will give some advices.

7-1 General troubleshooting

| Trouble | Cause | Remedy |
| --- | --- | --- |
| 1、Double images | Interpupillary distance is not correct | Readjust it |
|  | Diopter adjustment is not correct | Readjust it |
|  | Magnification of each eyepiece is not the same size | Mount the same size eyepiece |
| 2、Dirt appears in the field of view | Dirt on the specimen | Clean the specimen |
|  | Dirt on the surfaces of eyepiece | Clean the surface |
| 3、Image is not clear | Dirt on the surfaces of the objectives | Clean the objectives |
| 4、Image is not clear while the focus changing | Diopter adjustment is not correct | Readjust the diopter |
|  | Focus is not correct | Readjust the focus |
| 5、The focusing knob is not smooth | The focusing knob is too tight | Loosen it to a suitable position |
| 6、The image is obscure because of the head slipping down by itself during observation | The focusing knob is too loose | Tighten it to a suitable position |
| 7、Incision image appears in the field of view or of the video view | The pole is not in correct position | Pull or push it to the correct position |
| 8、The image on the monitor is not clear when the focusing knob is turned. | The focus of video is not correct | Readjust the focus of video to a correct position |
| 9、Eyes fell tired easily | Diopter adjustment is not correct | Adjust the diopter |
|  | Brightness of light is not correct | Adjust the brightness |
| 10、Bulb does not work when the switch is on | No power supply | Check the connection with the power supply |
|  | The bulb was not inserted correctly | Insert it correctly |
|  | Bulb is wrong | Replace with a new one |
| 11、Bulb is burned out suddenly | Use the wrong bulb | Replace with a correct one |
|  | The voltage is too high | Control the voltage Eg: use voltage regulator |
| 12、Brightness is not enough | Use a wrong bulb | Replace with a correct one |
|  | The voltage is too low | Increase the input voltage |
| 13、The bulb flickers or the brightness is unstable | The bulb will burn out soon | Replace with a new one |
|  | The bulb was not inserted correctly | Insert it correctly |

7-2 Video troubleshooting

| Trouble | Cause | Remedy |
| --- | --- | --- |
| 1、Incision image appears in the video view | The pole is not in correct position | Draw it to the correct position |
| 2、Dirt appears in the video view | Dirt on the specimen | Clean the specimen |
|  | Dirt on the surface of objective | Clean the surface |
| 3、Image is not clear while the focus changing | The image is not clear in the high magnification | Readjust the high magnification |
| 4、No image on the TV screen | The draw pole is not in correct position | Draw it to the correct position |
|  | Objective cover is not open | Open it |
|  | TV is not on Video channel | Choose the correct one |
| 5、No image on the Monitor | Connection is not correct | Reconnect the circuit |
|  | Objective cover is not open | Open it |
|  | The input signal does not accord with the signal be chosen on the Monitor | Choose the correct signal model |
| 6、The software run slowly or the window of the view does not come out | 12V DC power does not be connect | Connect the 12V DC power |
|  | No input signal of A/D board | Reconnect the C-Video or S-Video signal |
|  | The input signal does not accord with the signal which is chosen in the driver of the A/D board | Choose the correct signal model which match the input signal |
| 7、The image is not correct on the view window | The CCD model chosen in the driver of the A/D board does not accord with the real CCD | Choose the correct CCD model |